## calculator
U2000 / U8480 Series USB Power Sensors Uncertainty Calculator
Input | Value
Device Under Test (SWR) | 1.25
Sensor Model | U2000A
Frequency | 2 | GHz
Power | -13 | dBm |  | 5.0119e-05 | W
Number of Average Reading | 16
Average mode | Normal
Internal or External Zeroing (not applicable for U8480) | Internal zero
Measurement mode | Average only mode
Gate length (not required for Average only mode) | 2000 | us
Source of Uncertainty | Symbol | Value ± | Probability Distribution | Divisor | Result
Mismatch Gain Between Generator and Sensor | Mu | ¦Гg¦ = 0.111 | U-shape | 1.4142 | 0.6816%
¦Гs¦ = 0.087
Zero drift | D | 1.5e-09 | Rectangular | 1.7321 | 0.001728%
Power Sensor Calibration Factor Uncertainties | Kb | 2% | Gaussian | 2 | 1%
Power Sensor Linearity | Pl | 3% | Gaussian | 2 | 1.5%
Zero set | Zs | 1.2e-08 | Rectangular | 1.7321 | 0.01382%
Measurement noise | N | 1.5e-08 | Rectangular | 1.7321 | 0.01728%
Sensor Internal Calibration Uncertainties | Pcal | 0% | Gaussian | 2 | 0%
|  |  | Combined Uncertainty-RSSed | = | 1.927%
|  |  | K | = | 2
|  |  | Expanded Uncertainty | = | 3.855%
|  |  | Upper Limit Uncertainty | = | + 0.164dB
|  |  | Lower Limit Uncertainty | = | − 0.171dB
Revision 8
Date updated: 10 Feb 2014
© Agilent Technologies, Inc. 2008, 2011, 2012, 2013, 2014
The data used in this calculation is base on the specifications as in the U2000 series USB power sensors datasheet, literature number 5989-6278EN, and N8480 series datasheet, literature number 5991-1410EN).
This calculation is base on ISO Guide to the Expression of Uncertainty in Measurement, often referred to as the GUM.
For more info on power meter and power sensor measurement uncertainty equation and calculation base on GUM, please refer to Agilent Application Note 1449-3 "Fundamental of RF and Microwave Power Measurements (Part 3), literature number 5988-9215EN.

## Data (hidden)
Max SWR (25 +/-10C)
Frequency | U2000A | U2001A | U2002A | U2004A | U2000B | U2000H | U2001B | U2001H | U2002H | U8481A | U8485A | U8487A | U8488A
DC | # | # | # | # | # | # | # | # | # | 1.11 | 1.07 | # | #
1e-12 | # | # | # | # | # | # | # | # | # | 1.11 | 1.07 | # | #
9e-06 | # | # | # | 1.13 | # | # | # | # | # | 1.11 | 1.07 | # | # |  | 9kHz
0.01 | 1.15 | 1.15 | # | 1.13 | 1.12 | 1.15 | 1.12 | 1.15 | # | 1.37 | 1.33 | 1.35 | 1.06 |  | 10MHz
0.03 | 1.13 | 1.13 | # | 1.13 | 1.12 | 1.15 | 1.12 | 1.15 | # | 1.14 | 1.33 | 1.35 | 1.06 |  | 30MHz
0.05 | 1.13 | 1.13 | 1.13 | 1.13 | 1.12 | 1.15 | 1.12 | 1.15 | 1.15 | 1.08 | 1.08 | 1.08 | 1.06 |  | 50MHz
0.1 | 1.13 | 1.13 | 1.13 | 1.13 | 1.12 | 1.15 | 1.12 | 1.15 | 1.15 | 1.08 | 1.05 | 1.05 | 1.06 |  | 100MHz
2 | 1.19 | 1.19 | 1.19 | 1.19 | 1.17 | 1.15 | 1.17 | 1.15 | 1.15 | 1.16 | 1.14 | 1.1 | 1.06 |  | 2GHz
2.4 | 1.19 | 1.19 | 1.19 | 1.19 | 1.17 | 1.15 | 1.17 | 1.15 | 1.15 | 1.16 | 1.14 | 1.1 | 1.13 |  | 2.4GHz
6 | 1.19 | 1.19 | 1.19 | 1.19 | 1.17 | 1.15 | 1.17 | 1.15 | 1.15 | 1.16 | 1.14 | 1.1 | 1.13 |  | 6GHz
8 | 1.19 | 1.19 | 1.19 | 1.19 | 1.17 | 1.25 | # | # | 1.25 | 1.16 | 1.14 | 1.1 | 1.13 |  | 8GHz
12.4 | 1.19 | 1.19 | 1.19 | 1.19 | 1.24 | 1.28 | # | # | 1.28 | 1.23 | 1.19 | 1.16 | 1.14 |  | 12.4GHz
14 | 1.22 | # | 1.22 | # | 1.24 | 1.28 | # | # | 1.28 | 1.23 | 1.19 | 1.16 | 1.14 |  | 14GHz
16 | 1.26 | # | 1.26 | # | 1.24 | 1.28 | # | # | 1.28 | 1.23 | 1.19 | 1.16 | 1.14 |  | 16GHz
18 | 1.26 | # | 1.26 | # | 1.24 | 1.28 | # | # | 1.3 | 1.23 | 1.26 | 1.22 | 1.2 |  | 18GHz
19 | # | # | 1.3 | # | # | # | # | # | 1.3 | # | 1.26 | 1.22 | 1.2 |  | 19GHz
24 | # | # | 1.3 | # | # | # | # | # | 1.3 | # | 1.26 | 1.22 | 1.2 |  | 24GHz
26.5 | # | # | # | # | # | # | # | # | # | # | 1.37 | 1.3 | 1.25 |  | 26.5GHz
33 | # | # | # | # | # | # | # | # | # | # | 1.37 | 1.3 | 1.25 |  | 33GHz
40 | # | # | # | # | # | # | # | # | # | # | # | 1.34 | 1.42 |  | 40GHz
50 | # | # | # | # | # | # | # | # | # | # | # | 1.34 | 1.42 |  | 50GHz
67 | # | # | # | # | # | # | # | # | # | # | # | # | 1.36 |  | 67GHz
70 | # | # | # | # | # | # | # | # | # | # | # | # | 1.36 |  | 70GHz
Power Linearity (25+/-10C)
Power Level | U2000A | U2001A | U2002A | U2004A | U2000B | U2000H | U2001B | U2001H | U2002H | U8481A | U8485A | U8487A | U8488A
-60 | 3% | 3% | 3% | 3% | 3.5% | 4% | 3.5% | 4% | 4% | 0% | 0% | 0% | 0%
-35 | 3% | 3% | 3% | 3% | 3.5% | 4% | 3.5% | 4% | 4% | 0% | 0% | 0% | 0%
-1 | 3% | 3% | 3% | 3% | 3.5% | 4% | 3.5% | 4% | 4% | 0.5% | 0.5% | 0.5% | 0.5%
15 | 3% | 3% | 3% | 3% | 3.5% | 4% | 3.5% | 4% | 4% | 0.75% | 0.75% | 0.75% | 0.75%
20 | 3% | 3% | 3% | 3% | 3.5% | 4% | 3.5% | 4% | 4% | 0.75% | 0.75% | 0.75% | 0.75%
Internal zero
External zero
Zero Set (External) | (Note: For these parameters, there are overlapping power range. In this case, worst case spec is chosen for MU calculation).
| U2000A | U2001A | U2002A | U2004A |  |  | U2000B | U2001B |  |  | U2000H | U2001H | U2002H
-60 | 6e-10 | 6e-10 | 6e-10 | 6e-10 |  | -30 | 8e-07 | 8e-07 |  | -50 | 8e-09 | 8e-09 | 8e-09
-38 | 1.5e-09 | 1.5e-09 | 1.5e-09 | 1.5e-09 |  | -8 | 2e-06 | 2e-06 |  | -28 | 2e-08 | 2e-08 | 2e-08
-20 | 1e-08 | 1e-08 | 1e-08 | 1e-08 |  | 10 | 1e-05 | 1e-05 |  | -10 | 1e-07 | 1e-07 | 1e-07
-7.5 | 5e-07 | 5e-07 | 5e-07 | 5e-07 |  | 19 | 0.001 | 0.001 |  | -1 | 2e-05 | 2e-05 | 2e-05
-4 | 1e-06 | 1e-06 | 1e-06 | 1e-06 |  | 23 | 0.002 | 0.002 |  | 3 | 3e-05 | 3e-05 | 3e-05
10 | 5e-06 | 5e-06 | 5e-06 | 5e-06 |  | 44 | 0.002 | 0.002 |  | 20 | 6e-05 | 6e-05 | 6e-05
20 | 5e-06 | 5e-06 | 5e-06 | 5e-06 | [5 empty] | 30 | 6e-05 | 6e-05 | 6e-05
Zero Set (Internal) | U2000A | U2001A | U2002A | U2004A |  |  | U2000B | U2001B |  |  | U2000H | U2001H | U2002H
-60 | 1.5e-09 | 1.5e-09 | 1.5e-09 | 2.8e-09 |  | -30 | 1.8e-06 | 1.8e-06 |  | -50 | 1.5e-09 | 1.5e-09 | 1.5e-09
-38 | 2e-09 | 2e-09 | 2e-09 | 3e-09 |  | -8 | 2e-06 | 2e-06 |  | -28 | 2e-08 | 2e-08 | 2e-08
-20 | 1.2e-08 | 1.2e-08 | 1.2e-08 | 1.2e-08 |  | 10 | 1.2e-05 | 1.2e-05 |  | -10 | 1.2e-07 | 1.2e-07 | 1.2e-07
-7.5 | 2e-06 | 2e-06 | 2e-06 | 2e-06 |  | 19 | 0.002 | 0.002 |  | -1 | 2e-05 | 2e-05 | 2e-05
-4 | 4e-06 | 4e-06 | 4e-06 | 4e-06 |  | 23 | 0.004 | 0.004 |  | 3 | 4e-05 | 4e-05 | 4e-05
10 | 6e-06 | 6e-06 | 6e-06 | 6e-06 |  | 44 | 0.004 | 0.004 |  | 20 | 6e-05 | 6e-05 | 6e-05
20 | 6e-06 | 6e-06 | 6e-06 | 6e-06 | [5 empty] | 30 | 6e-05 | 6e-05 | 6e-05
Zero Drift | U2000A | U2001A | U2002A | U2004A |  |  | U2000B | U2001B |  |  | U2000H | U2001H | U2002H
-60 | 2e-10 | 2e-10 | 2e-10 | 2e-10 |  | -30 | 2e-07 | 2e-07 |  | -50 | 2e-09 | 2e-09 | 2e-09
-38 | 4e-10 | 4e-10 | 4e-10 | 4e-10 |  | -8 | 4e-07 | 4e-07 |  | -28 | 4e-09 | 4e-09 | 4e-09
-20 | 1.5e-09 | 1.5e-09 | 1.5e-09 | 1.5e-09 |  | 10 | 1.5e-06 | 1.5e-06 |  | -10 | 1.5e-08 | 1.5e-08 | 1.5e-08
-7.5 | 5e-08 | 5e-08 | 5e-08 | 5e-08 |  | 19 | 5e-08 | 5e-08 |  | -1 | 5e-07 | 5e-07 | 5e-07
-4 | 5e-07 | 5e-07 | 5e-07 | 5e-07 |  | 23 | 0.0005 | 0.0005 |  | 3 | 5e-06 | 5e-06 | 5e-06
10 | 2e-06 | 2e-06 | 2e-06 | 2e-06 |  | 44 | 0.0005 | 0.0005 |  | 20 | 2e-05 | 2e-05 | 2e-05
20 | 2e-06 | 2e-06 | 2e-06 | 2e-06 | [5 empty] | 30 | 2e-05 | 2e-05 | 2e-05
Sensor noise | U2000A | U2001A | U2002A | U2004A |  |  | U2000B | U2001B |  |  | U2000H | U2001H | U2002H
-60 | 1e-09 | 1e-09 | 1e-09 | 1e-09 |  | -30 | 1e-06 | 1e-06 |  | -50 | 1e-08 | 1e-08 | 1e-08
-38 | 1.5e-09 | 1.5e-09 | 1.5e-09 | 1.5e-09 |  | -8 | 1.5e-06 | 1.5e-06 |  | -28 | 1.5e-08 | 1.5e-08 | 1.5e-08
-20 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 |  | 10 | 1.5e-05 | 1.5e-05 |  | -10 | 1.5e-07 | 1.5e-07 | 1.5e-07
-7.5 | 6.5e-07 | 6.5e-07 | 6.5e-07 | 6.5e-07 |  | 19 | 0.00065 | 0.00065 |  | -1 | 6.5e-06 | 6.5e-06 | 6.5e-06
-4 | 1e-06 | 1e-06 | 1e-06 | 1e-06 |  | 23 | 0.001 | 0.001 |  | 3 | 1e-05 | 1e-05 | 1e-05
10 | 1e-05 | 1e-05 | 1e-05 | 1e-05 |  | 44 | 0.001 | 0.001 |  | 20 | 0.0001 | 0.0001 | 0.0001
20 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | [5 empty] | 30 | 0.0001 | 0.0001 | 0.0001
Normal | x2 | Fast
Noise multiplier | U2000 |  |  |  |  | U2000
Number of avg | Normal |  |  |  | Number of avg | x2
1 | 2 |  |  |  | 1 | 2.7
2 | 1.8 |  |  |  | 2 | 2.4
4 | 1.7 |  |  |  | 4 | 2
8 | 1.5 |  |  |  | 8 | 1.6
16 | 1 |  |  |  | 16 | 1
32 | 0.95 |  |  |  | 32 | 0.91
64 | 0.74 |  |  |  | 64 | 0.78
128 | 0.55 |  |  |  | 128 | 0.53
256 | 0.39 |  |  |  | 256 | 0.34
512 | 0.29 |  |  |  | 512 | 0.29
1024 | 0.21 |  |  |  | 1024 | 0.2
U2000 multiplier | 1
Cal factor uncertainties
Frequency | U2000A | U2001A | U2002A | U2004A | U2000B | U2000H | U2001B | U2001H | U2002H | U8481A | U8485A | U8487A | U8488A
DC | # | # | # | # | # | # | # | # | # | 2.63% | 2.37% | # | #
1e-13 | # | # | # | # | # | # | # | # | # | 2.63% | 2.37% | # | #
9e-06 | # | # | # | 1.8% | # | # | # | # | # | 2.63% | 2.37% | # | # | 9kHz
0.01 | 1.8% | 1.8% | # | 1.8% | 1.8% | 2% | 1.8% | 2% | # | 1.05% | 1.5% | 1.79% | 2.04% | 10Mhz
0.03 | 1.6% | 1.6% | # | 1.8% | 1.8% | 2% | 1.8% | 2% | # | 0.85% | 1.37% | 1.78% | 2.04% | 30MHz
0.05 | 1.6% | 1.6% | 2% | 1.8% | 1.8% | 2% | 1.8% | 2% | 2.5% | 0.85% | 1.37% | 1.78% | 1.94% | 50MHz
0.1 | 1.6% | 1.6% | 2% | 1.8% | 1.8% | 2% | 1.8% | 2% | 2.5% | 0.85% | 1.37% | 1.78% | 1.98% | 100MHz
0.5 | 1.6% | 1.6% | 2% | 1.8% | 1.8% | 2% | 1.8% | 2% | 2.5% | 0.78% | 1.26% | 1.79% | 1.98% | 500MHz
1.2 | 1.6% | 1.6% | 2% | 1.8% | 1.8% | 2% | 1.8% | 2% | 2.5% | 0.91% | 1.35% | 1.82% | 1.98% | 1.2GHz
2 | 2% | 2% | 2.5% | 1.8% | 2% | 2% | 2% | 2% | 2.5% | 0.91% | 1.35% | 1.82% | 2.13% | 2GHz
6 | 2% | 2% | 2.5% | 1.8% | 2% | 2% | 2% | 2% | 2.5% | 1.26% | 1.66% | 1.88% | 2.13% | 6GHz
8 | 2% | 2% | 2.5% | 1.8% | 2% | 2% | # | # | 2.5% | 1.26% | 1.66% | 1.88% | 2.13% | 8GHz
12.4 | 2% | 2% | 2.5% | 1.8% | 2.2% | 2.2% | # | # | 2.7% | 1.26% | 1.66% | 1.88% | 2.24% | 12.4GHz
14 | 2.2% | # | 2.7% | # | 2.2% | 2.2% | # | # | 2.7% | 1.59% | 1.83% | 1.9% | 2.24% | 14GHZ
16 | 2.2% | # | 2.7% | # | 2.2% | 2.2% | # | # | 2.7% | 1.59% | 1.83% | 1.9% | 2.24% | 16GHz
18 | 2.2% | # | 3% | # | 2.2% | 2.2% | # | # | 3% | 1.59% | 2.67% | 2.09% | 2.52% | 18GHz
24 | # | # | 3% | # | # | # | # | # | 3% | # | 2.67% | 2.09% | 2.52% | 24GHz
26.5 | # | # | # | # | # | # | # | # | # | # | 3.32% | 2.66% | 4.66% | 26.5GHz
33 | # | # | # | # | # | # | # | # | # | # | 3.32% | 2.66% | 4.66% | 33GHz
40 | # | # | # | # | # | # | # | # | # | # | # | 3.73% | 4.66% | 40GHz
45 | # | # | # | # | # | # | # | # | # | # | # | 4.68% | 4.66% | 45GHz
50 | # | # | # | # | # | # | # | # | # | # | # | 4.68% | 5.14% | 50GHz
67 | # | # | # | # | # | # | # | # | # | # | # | # | 5.7% | 67GHz
70 | # | # | # | # | # | # | # | # | # | # | # | # | 5.7% | 70GHz
rho DUT | 0.11111
rho sensor | 0.086758 | 1.19
Multiplier | 1
Rectangular | 1.7321
Gaussian | 2
U-shape | 1.4142
U2000A |  |  | U2000A
U2001A |  |  | U2001A
U2002A |  |  | U2002A
U2004A |  |  | U2004A
U2000B |  |  | U2000B
U2000H |  |  | U2000H
U2001B |  |  | U2001B
U2001H |  |  | U2001H |  |  | Average only mode
U2002H |  |  | U2002H |  |  | Time-gated mode (normal mode)
U8481A |  |  | U8481A
U8485A |  |  | U8485A
U8487A |  |  | U8487A
U8488A |  |  | U8488A
Zero Set | Average only mode | Time-gated mode (normal mode) |  |  |  | Zero Drift
External zero | 1e-08 | 4.8e-07 |  |  |  | Average only mode | 1.5e-09
Internal zero | 1.2e-08 | 5.3e-07 |  |  |  | Time-gated mode (normal mode) | 2.3e-07
Time-gated mode (Normal Mode) Related Specifications
Power Linearity (25+/-10C)
| U2000A | U2001A | U2002A | U2004A | U2000B | U2000H | U2001B | U2001H | U2002H
| 4% | 4% | 4% | # | 4.5% | 5% | 4.5% | 5% | 5%
Zero Set (External) | (Note: For these parameters, there are overlapping power range. In this case, worst case spec is chosen for MU calculation).
| U2000A | U2001A | U2002A | U2004A |  |  | U2000B | U2001B |  |  | U2000H | U2001H | U2002H
-60 | # | # | # | # |  | -30 | # | # |  | -50 | # | # | #
-38 | 4.3e-08 | 4.3e-08 | 4.3e-08 | # |  | -8 | 4.3e-05 | 4.3e-05 |  | -28 | 5e-07 | 5e-07 | 5e-07
-20 | 4.8e-07 | 4.8e-07 | 4.8e-07 | # |  | 10 | 0.00048 | 0.00048 |  | -10 | 4.8e-06 | 4.8e-06 | 4.8e-06
-7.5 | 2.7e-05 | 2.7e-05 | 2.7e-05 | # |  | 19 | 0.027 | 0.027 |  | -1 | 0.00027 | 0.00027 | 0.00027
-4 | 3e-05 | 3e-05 | 3e-05 | # |  | 23 | 0.034 | 0.034 |  | 3 | 0.0003 | 0.0003 | 0.0003
10 | 0.0002 | 0.0002 | 0.0002 | # |  | 44 | 0.034 | 0.034 |  | 20 | 0.0028 | 0.0028 | 0.0028
20 | 0.0002 | 0.0002 | 0.0002 | # | [5 empty] | 30 | 0.0028 | 0.0028 | 0.0028
Zero Set (Internal) | (Note: For these parameters, there are overlapping power range. In this case, worst case spec is chosen for MU calculation).
| U2000A | U2001A | U2002A | U2004A |  |  | U2000B | U2001B |  |  | U2000H | U2001H | U2002H
-60 | # | # | # | # |  | -30 | # | # |  | -50 | # | # | #
-38 | 4.7e-08 | 4.7e-08 | 4.7e-08 | # |  | -8 | 4.7e-05 | 4.7e-05 |  | -28 | 7.3e-07 | 7.3e-07 | 7.3e-07
-20 | 5.3e-07 | 5.3e-07 | 5.3e-07 | # |  | 10 | 0.00053 | 0.00053 |  | -10 | 5.3e-06 | 5.3e-06 | 5.3e-06
-7.5 | 3e-05 | 3e-05 | 3e-05 | # |  | 19 | 0.03 | 0.03 |  | -1 | 0.00033 | 0.00033 | 0.00033
-4 | 3.2e-05 | 3.2e-05 | 3.2e-05 | # |  | 23 | 0.032 | 0.032 |  | 3 | 0.00044 | 0.00044 | 0.00044
10 | 0.00027 | 0.00027 | 0.00027 | # |  | 44 | 0.032 | 0.032 |  | 20 | 0.0039 | 0.0039 | 0.0039
20 | 0.00027 | 0.00027 | 0.00027 | # | [5 empty] | 30 | 0.0039 | 0.0039 | 0.0039
Zero Drift | U2000A | U2001A | U2002A | U2004A |  |  | U2000B | U2001B |  |  | U2000H | U2001H | U2002H
-60 | # | # | # | # |  |  | # | # |  |  | # | # | #
-38 | 2.5e-08 | 2.5e-08 | 2.5e-08 | 2.5e-08 |  | -8 | 2.5e-05 | 2.5e-05 |  | -28 | 3e-07 | 3e-07 | 3e-07
-20 | 2.3e-07 | 2.3e-07 | 2.3e-07 | 2.3e-07 |  | 10 | 0.00023 | 0.00023 |  | -10 | 3e-06 | 3e-06 | 3e-06
-7.5 | 1.9e-05 | 1.9e-05 | 1.9e-05 | 1.9e-05 |  | 19 | 0.019 | 0.019 |  | -1 | 0.00019 | 0.00019 | 0.00019
-4 | 2.4e-05 | 2.4e-05 | 2.4e-05 | 2.4e-05 |  | 23 | 0.024 | 0.024 |  | 3 | 0.0003 | 0.0003 | 0.0003
10 | 0.00011 | 0.00011 | 0.00011 | 0.00011 |  | 44 | 0.024 | 0.024 |  | 20 | 0.0011 | 0.0011 | 0.0011
20 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | [5 empty] | 30 | 0.0011 | 0.0011 | 0.0011
Measurement noise | U2000A | U2001A | U2002A | U2004A |  |  | U2000B | U2001B |  |  | U2000H | U2001H | U2002H
-60 | # | # | # | # |  | -30 | # | # |  | -50 | # | # | #
-38 | 2.8e-08 | 2.8e-08 | 2.8e-08 | 2.8e-08 |  | -8 | 2.8e-05 | 2.8e-05 |  | -28 | 3.1e-07 | 3.1e-07 | 3.1e-07
-20 | 3e-07 | 3e-07 | 3e-07 | 3e-07 |  | 10 | 0.0003 | 0.0003 |  | -10 | 5e-06 | 5e-06 | 5e-06
-7.5 | 2e-05 | 2e-05 | 2e-05 | 2e-05 |  | 19 | 0.02 | 0.02 |  | -1 | 0.00023 | 0.00023 | 0.00023
-4 | 2.1e-05 | 2.1e-05 | 2.1e-05 | 2.1e-05 |  | 23 | 0.021 | 0.021 |  | 3 | 0.00026 | 0.00026 | 0.00026
10 | 0.00018 | 0.00018 | 0.00018 | 0.00018 |  | 44 | 0.021 | 0.021 |  | 20 | 0.0028 | 0.0028 | 0.0028
20 | 0.00018 | 0.00018 | 0.00018 | 0.00018 | [5 empty] | 30 | 0.0028 | 0.0028 | 0.0028
Prelim Spec
Noise Persample | U2000A | U2001A | U2002A | U2004A |  |  | U2000B | U2001B |  |  | U2000H | U2001H | U2002H
-60 | # | # | # | # |  | -30 | # | # |  | -50 | # | # | #
-38 | 9e-08 | 9e-08 | 9e-08 | 9e-08 |  | -8 | 9e-05 | 9e-05 |  | -28 | 0.0009 | 0.0009 | 0.0009
-20 | 1e-06 | 1e-06 | 1e-06 | 1e-06 |  | 10 | 0.001 | 0.001 |  | -10 | 1e-05 | 1e-05 | 1e-05
-7.5 | 5.5e-05 | 5.5e-05 | 5.5e-05 | 5.5e-05 |  | 19 | 0.055 | 0.055 |  | -1 | 0.00055 | 0.00055 | 0.00055
-4 | 8.5e-05 | 8.5e-05 | 8.5e-05 | 8.5e-05 |  | 23 | 0.085 | 0.085 |  | 3 | 0.00085 | 0.00085 | 0.00085
10 | 0.00055 | 0.00055 | 0.00055 | 0.00055 |  | 44 | 0.085 | 0.085 |  | 20 | 0.0055 | 0.0055 | 0.0055
20 | 0.00055 | 0.00055 | 0.00055 | 0.00055 | [5 empty] | 30 | 0.0055 | 0.0055 | 0.0055
Measurement noise
Average only mode | 1.5e-08 | 1.5e-08
Time gated mode | 3e-07
Time-gated mode <2.73us | 2.5e-07 | 1e-06
Time-gated mode >2.73us | 4.8009e-08
Measurement noise (normal mode) | 3e-07
U8480 series
Zero set | 2.5e-08
Zero drift | 5.5e-09
Measurement noise | 4.5e-08
U8480 Multiplier | 1
| U8481/85A | U8481/85A | U8481/5A
Number of avg | Normal | x2 | Fast
1 | 3.17 | 4.55 | 46.88
2 | 2.62 | 3.76 | 33.06
4 | 2.02 | 3 | 24
8 | 1.54 | 2.25 | 17.19
16 | 1 | 1.59 | 12.24
32 | 0.82 | 1 | 8.39
64 | 0.6 | 0.85 | 4.93
128 | 0.5 | 0.63 | 4.11
256 | 0.37 | 0.47 | 2.48
512 | 0.27 | 0.42 | 1
1024 | 0.15 | 0.23 | 0.83
Note on K103: sookhua: 25+/-3degC
Note on L103: sookhua: 25+/-3degC
Note on M103: sookhua: 25+/-3C
Note on N103: sookhua: 25+/-3C

## Revision history
Revision 1 | New Creation
Revision 2 | Added U2000/1/2H and U2000/1B
Revision 3 | Added Time-gated (normal mode) measurement uncertainty analysis
Revision 4 | Minor update on spec change.
Revision 5 | Corrected error in mismatch uncertainty calculation
Revision 6 | Corrected error in mismatch uncertainty calculation. Updated power range for zero set/drift and measurement noise.
Revision 7 | Added U8481A and U8485A
Revision 8 | Added U8487A and U8488A