## Лист1
Standard type | Model | S/N | Reference value | Units | Accuracy limit | Actual value | Error | Uncertainty | Annual drift (ʋ, ppm) | TC(ppm/°C) |  | Verification date | Verification lab | Previous value | Previous verification date | Previous value | Previous verification date
[8 empty] | (ppm) |  | α | β
Capacitance | P597/6 | 182 | 400 | pF | ±0,05% | 399.883 | -0.02925% | 200 | -60.9494 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 399.918 | 2021-10-06
Capacitance | P597/6 | 5260 | 400 | pF | ±0,05% | 399.927 | -0.01825% | 200 |  |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva
Capacitance | P597/3 | 2267 | 100 | pF | ±0,05% | 99.914 | -0.086% | 200 | -856.7748 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 100.037 | 2021-10-06
Capacitance | P597/3 | 5283 | 100 | pF | ±0,05% | 99.869 | -0.131% | 200 |  |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva
Capacitance | КМЕ-101 | 1249-1987 | 10 | pF | ±0,02% | 10.0018 | 0.018% | 200 | 0 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 10.0018 | 2021-08-31
Capacitance | КМЕ-101 | 925-1982 | 10 | pF | ±0,02% | 10.0014 | 0.014% | 200 | -18.5818 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 10.0017 | 2021-09-10
Capacitance | КМЕ-11 | 3290-1981 | 1 | pF | ±0,02% | 1.0002 | 0.022% | 200 | 14.2857 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 1.0002 | 2021-10-19
Capacitance | КМЕ-11 | 4375-1982 | 1 | pF | ±0,02% | 1.0001 | 0.009% | 200 | -6.5179 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 1.0001 | 2021-08-31 | 1.0001 | 1987-01-04
Capacitance | КМЕ-11 | 5175-1984 | 1 | pF | ±0,02% | 1.0001 | 0.014% | 200 | 13.2486 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 1.0001 | 2021-09-09
Capacitance | КМЕ-11 | 3291-1981 | 0.5 | pF | ±0,02% | 0.50004 | 0.0076% | 200 | -68.5714 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 0.50009 | 2021-10-19
Capacitance | КМЕ-11 | 4376-1982 | 0.5 | pF | ±0,02% | 0.5 | 0.0008% | 200 | 7.8214 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 0.5 | 2021-08-31 | 0.50004 | 1987-01-04
Capacitance | КМЕ-11 | 5176-1984 | 0.5 | pF | ±0,02% | 0.50008 | 0.0164% | 200 | -9.274 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 0.50009 | 2021-09-09
Capacitance | КМЕ-11 | 3292-1981 | 0.1 | pF | ±0,02% | 0.10003 | 0.034% | 200 | -100 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 0.10005 | 2021-10-19
Capacitance | КМЕ-11 | 4377-1982 | 0.1 | pF | ±0,02% | 0.09999 | -0.01% | 200 | 0 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 0.09999 | 2021-08-31 | 0.09999 | 1987-01-04
Capacitance | КМЕ-11 | 5177-1984 | 0.1 | pF | ±0,02% | 0.10004 | 0.041% | 200 | 331.216 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 0.099991 | 2021-09-09
Capacitance | КМЕ-11 | 3293-1981 | 0.01 | pF | ±0,05% | 0.010002 | 0.023% | 500 | -192.8571 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 0.010005 | 2021-10-19
Capacitance | КМЕ-11 | 4378-1982 | 0.01 | pF | ±0,05% | 0.010003 | 0.03% | 500 | 65.1786 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 0.010002 | 2021-08-31 | 0.009997 | 1987-01-04
Capacitance | КМЕ-11 | 5178-1984 | 0.01 | pF | ±0,05% | 0.009999 | -0.01% | 500 | -132.4864 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 0.010001 | 2021-09-09
Capacitance | КМЕ-11 | 3294-1981 | 0.001 | pF | ±0,3% | 0.0010011 | 0.11% | 3000 | 1071.4286 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 0.0009996 | 2021-10-19
Capacitance | КМЕ-11 | 4379-1982 | 0.001 | pF | ±0,3% | 0.0010008 | 0.08% | 3000 | 130.3571 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 0.0010006 | 2021-08-31 | 0.001001 | 1987-01-04
Capacitance | КМЕ-11 | 5179-1984 | 0.001 | pF | ±0,3% | 0.0010011 | 0.11% | 3000 | 953.902 |  |  | 2023-03-14 | VNIIM D.I.Mendeleeva | 0.00099966 | 2021-09-09
Resistance | P4030 | 1285 | 1000 | MΩ | ±0,01% | 1000.043 | 0.0043% | 4 | 1.901 | 8.0364 | -0.0026143 | 2023-03-22 | VNIIM D.I.Mendeleeva | 1000.041 | 2022-03-03
Resistance | P4033 | 799 | 100 | MΩ | ±0,005% | 99.9936 | -0.0064% | 3 | -3.8021 | 13.6291 | -0.0096847 | 2023-03-22 | VNIIM D.I.Mendeleeva | 99.994 | 2022-03-03
Resistance | P4023 | 624 | 10 | MΩ | ±0,005% | 9.9992 | -0.008% | 2 | -1.901 | 9.8302 | -0.0079505 | 2023-03-22 | VNIIM D.I.Mendeleeva | 9.9992 | 2022-03-03
Resistance | P4023 | 0049 | 10 | MΩ | ±0,005% | 9.9993 | -0.0072% | 2 | 6.6536 | 3.374 | -0.082413 | 2023-03-22 | VNIIM D.I.Mendeleeva | 9.9992 | 2022-03-03
Resistance | P4013 | 153 | 1 | MΩ | ±0,005% | 0.99999 | -0.0008% | 2 | 0.95052 | -0.40938 | -0.056922 | 2023-03-22 | VNIIM D.I.Mendeleeva | 0.99999 | 2022-03-03
Resistance | P4015 | 301 | 100 | kΩ | ±0,005% | 100.0008 | 0.0008% | 2 | 0 | -4.6012 | 0.063496 | 2023-03-22 | VNIIM D.I.Mendeleeva | 100.0008 | 2022-03-03
Resistance | P331 | 196317-1971 | 10 | kΩ | ±0,01% | 10 | -0.0002% | 2 |  | 7.6157 | -0.57648 | 2023-03-21 | VNIIM D.I.Mendeleeva
Resistance | P331 | 079600-1981 | 1 | kΩ | ±0,01% | 1.0001 | 0.0088% | 2 |  | 10.2251 | -0.54592 | 2023-03-21 | VNIIM D.I.Mendeleeva
Resistance | P331 | 354560-1977 | 100 | Ω | ±0,01% | 99.9944 | -0.0056% | 2 |  | 10.6 | -0.84 | 2023-03-21 | VNIIM D.I.Mendeleeva
Resistance | P321 | 144912-1988 | 10 | Ω | ±0,01% | 9.9999 | -0.001% | 2 |  | 1.1 | -0.51 | 2023-03-21 | VNIIM D.I.Mendeleeva
Resistance | P321 | 142131-1967 | 1 | Ω | ±0,01% | 0.99995 | -0.0049% | 2 |  | -5.6 | -0.45 | 2023-03-21 | VNIIM D.I.Mendeleeva
Resistance | P321 | 137452-1990 | 100 | mΩ | ±0,01% | 99.999 | -0.001% | 2 |  | 5.4 | -0.6 | 2023-03-21 | VNIIM D.I.Mendeleeva
Resistance | P310 | 178073-1990 | 10 | mΩ | ±0,01% | 10.0014 | 0.0141% | 2 |  |  |  | 2023-03-21 | VNIIM D.I.Mendeleeva
Resistance | P310 | 056502-1981 | 1 | mΩ | ±0,01% | 1 | 0.0028% | 2 |  |  |  | 2023-03-21 | VNIIM D.I.Mendeleeva
Voltage | B1-30 | 0215 | 6.4 | V | ±0,005% | 6.5018 |  | 4 |  |  |  | 2023-10-30 | VNIIFTRI
Voltage | B1-30 | 0049 | 6.4 | V | ±0,005% | 6.4595 |  | 4 |  |  |  | 2023-10-30 | VNIIFTRI
Temperature | LTA/2Б-Э-Э | 873219 | -50…+200 | °C | ±0,02°C | [6 empty] | 2023-06-13 | Termex JSC